Labnovation Technologies, Inc.

Tel.: +86 755 86368328 (Ext.830)

Fax: +86 755 86368318

http://www.labnovation.com/

Address: 1F, W2-A Building, Keji 2nd Road South, Hi-Tech Industrial Park, Nanshan District, Shenzhen 518057, China

LD-560 communication protocol

This system by default uses the network interface to carry on the two-way communication.

Please use general Internet cable to connect the instrument with PC

PC settings(TCP /IPv4 ) :IP (192.168.0.101)，subnet mask(255.255.255.0)，default gateway(192.168.0.1)

LIS system settings: server IP(192.168.0.108), server port (8000)

Instrument settings : IP address (192.168.0.108), Mask Address (255.255.255.0) ,

Gateway (192.168.0.1)

2、data transmission format：

<TRANSMIT><M>LD560|LD560-001<I>sample|2018-03-15 22:34:54|3105|10|1|10|0</I><R>HbA1a|1.04HbA1b|1.01HbF|1.5

L-A1C|1.0HbA1c|7.19HbA0|92eAG|4.5</R></M></TRANSMIT>

| Field | Description and meaning |
| --- | --- |
| <TRANSMIT> | Message start symbol |
| </TRANSMIT> | Message end symbol |
| <M>、¦、</M> | Separator and tag of the instrument model name |
| LD560¦LD560-001 | Instrument model name and sequential number |
| <I>、¦、</I> | Separator and tag of the sample information |
| Sample | sample |
| 0,1,2,3,4 | 0: Whole blood 1: QC mode` 2: Calibration 3: Pre-diluted 4:emergency mode |
| 2018-03-15 22:34:54 | The time and date when each test finished. |
| 3105¦10¦1¦10¦0 | 3105: Sequential No 10: sample ID 1: rack No 10: tube No 0: sample type |
| <R>、¦、</R> | Separator and tag of the test results |
| HbA1a：1.04 | The test result of HbA1a |
| HbA1b：1.01 | The test result of HbA1b |
| HbF：1.5 | The test result of HbF |
| L-A1C：1.0 | The test result of L-A1C |
| HbA1c: 7.19 | The test result of HbA1c |
| HbA0: 92 | The test result of HbA0 |
| eAG : 4.5 | The test result of eAG |